LCQ Guide – Basic deconvolution using ProMass Application

EJL 3.1.07

This guide explains how to perform mass deconvolution on a .raw spectrum obtained on the Finnigan LCQ ESI mass spectrometer using the ProMass standalone application.  It assumes basic user knowledge of instrument and software operation.

- Accumulate the convoluted mass spectrum as a .raw file.

- Open the .raw file in Qual Browser.

- Select the chromatogram and mass ranges that you would like to deconvolute.

- Right-click in the mass spectrum window, and in the pop-up menu select Export->Clipboard (Exact Mass).  Depending on file size this process could take up to a minute.

- Once the copy has completed, Qual Browser can be closed.  Open ProMass for XCalibur; it should have a shortcut on the desktop.

- Click Build Params on the main menu screen.  This takes you to the page of options for customizing deconvolution.

- On the top of the window click the Basic Deconvolution tab to switch to basic mode.  Advanced mode can be used but it will not be included in this guide.

- Follow the steps outlined on the parameter page (1-4).  Clicking the “Paste and Process Clipboard Spectrum” button (step 3) will initiate the deconvolution process.  This should open a DOS window that will quickly scroll through the various calculations involved.

- Upon completion of the deconvolution process the program will open a web browser window with the formatted results.  This includes graphs of the convoluted and deconvoluted spectra and a detail of the base peak in the deconvolution.

- The results displayed in the web browser are saved as a collection of image/text files in C:\Program Files\ProMassXcali\results\promass_results.  The filenames are based on the original .raw filename, and each file contains the following:

  - (Filename).cent.txt – deconvolution labeled-peak mass lists.

  - (Filename).dec – full numerical data for convoluted and deconvoluted spectra.  This file can be opened and manipulated graphically using ProMassDecView (shortcut on desktop).  Alternately it can be opened as a text file and the raw numerical data can be viewed/extracted.

  - (Filename).html – html document used to display the results via web browser.

  - (Filename).log.txt – text file displaying deconvolution parameters used.

  - (Filename).dec.jpeg – jpeg image of the full deconvoluted spectrum.

  - (Filename).jpeg – jpeg image of the original convoluted spectrum input.

  - (Filename).zdec.jpeg – jpeg image of the base peak detail in the deconvoluted spectrum.