SPECIFICATION

MODEL 2554-2

Voltage and Current:	0 to +30kV @ 0 to 400(Adc

Polarity:	User selectable via remote input logic signal compatible with HCMOS, open collector and relay logic. If polarity reversal is commanded during HIGH VOLTAGE operation, the unit will automatically turn off, wait for discharge, reverse polarity and then resume HIGH VOLTAGE operation at the new polarity; at the same output voltage magnitude. The polarity reversing time duration will be less than 3.5 seconds, measured from the time of switching to 90% of final voltage.

Power:					+24Vdc (+5%/-2%) @ 2Adc maximum (1.5Adc typical).

Voltage Regulation: 			0.001% for input change of 1 volt.

0.001 % for 100(A to full load.

Current Regulation:			+0.05% for (+5%/2%) Input line change.

+0.1 % for 0 to maximum rated output voltage.

Ripple:					0.001% peak-to-peak at maximum output.

Temperature Coefficient (0 to 50(C):	100ppm of maximum per (C.

Stability (after 1/2 hr warm-up):	0.01% per hour; 0.02% per 8 hours.

Energy Storage:			Less than 0.75 Joule.

Discharge Time:	The power supply will discharge with an initial time constant of less than 1 sec. (With 1000 megohm load, discharge to 10% of VOUTPUT in <2.25 seconds).

Local Output Controls:	Two independent multi-turn potentiometers for voltage and current control.

Remote Programming:	Two 0 to + 10Vdc programming inputs can be used to independently control the output high voltage and current from 0 to their maximum ratings. Accuracy is +(0.25% of setting + 0.25% of maximum). The programming input impedance is 10 megohm.

Voltage Monitor:	0 to 10Vdc +(0.25% of reading + 0.25% of max.) for 0 to 30 kV output. The monitor output impedance is 1 kilohms

Current Monitor:	0 to +10Vdc +0.25% of reading + 0.25% of max.) for 0 to 400uA output. The monitor output impedance is 1 kilohms.

Interlock:	A remote safety, ENABLE, input is provided. The remote ENABLE must be connected to COMMON to turn ON the high voltage output.

Protection:				Arc and short circuit protection, self-restoring.

Size:					Option -1: 5"W x 5-1/2"H x 8-1/2"L (127 x 140 x 216mm)

Option -2: 5"W x 3"H x 15"L (127 x 77 381 mm)

Weight:				12 Ibs (5.5kg)

High Voltage Connector:	Alden 8101QFP.  A three meter long mating high voltage cable assembly, Bertan P/N B200J118, is provided with each unit.

Power Input and l/O Connector: 	Standard 15 pin "D" type connector provides the functions as listed in the chart below. A mating connector is provided.

PIN#	FUNCTION

1	VOUTPUT Local Control (Jumper Pin 1 to Pin 2

2	VOUTPUT Remote Program (0 to +10Vdc Input)

3	IOUTPUT Local Control (Jumper Pin 3 to Pin 4)

4	IOUTPUT Remote Program (0 to +10Vdc Input)

5	VOUTPUT Monitor (0 to +10Vdc Output)

6 	Polarity Control (POSITVE/NEGATIVE)

7	Common

8	+24Vdc Return (COMMON)

9	N.C.

10	N.C.

11	N.C.

12	Interlock/ENABLE Logic Input

13	IOUTPUT Monitor (0 to +1 0Vdc Output)

14	+10Vdc Reference Output

15	+24Vdc Input Power

NOTE: The / before the function indicates that the function is active when low.

ACCESSORIES: 413043 accessory kit

Rev. 931224 BM/EM/PCC

CHANGE SUMMARY:

- Changed HV cable assembly P/N from 8101QM3NLB40 to B200J118.

Rev. 020731 BM/EM/PCC/JT